User Manual

Indigo-imaging Technical CO.,LTD

version 1.3

2021.4

The trademarks and names of the hardware and software products mentioned in this manual are the property of their respective owners.

The copyright of this manual belongs to Indigo-imaging Co., Ltd. All rights reserved. No formal approval of the Company, any organization or individual shall in any means and formats of this manual may be reproduced or transmitted.

The contents of this manual are subject to change without notice.

Copyright © 2016 Indigo-imaging Co., Ltd. All rights reserved.

Tel: 0591-83588265

http: //www.indigo-imaging.com

E-mail: support@indigo-imaging.com

Contents

1 Overview	1

1.1 Manual Instructions	1

1.2 Naming Rules	1

1.3 Packing List	1

1.4 Safety Precautions	2

1.5 Dimensions	2

1.5.1 Camera Dimensions	2

1.5.2 PCB Dimensions	3

1.6 Camera Features	4

1.7  Technical Specifications	5

1.8 Spectral Response	6

1.9 Chip Temperature	7

1.9.1 Working Temperature	7

1.9.2 Relationship between dark current and temperature	7

1.10 Applications	8

2 Physical Interface and I/O Trigger Connector	8

2.1 Camera Interface Functions	8

2.2 The I/O Trigger Pins and Connector Interface Functions	9

2.2.1 The Pins Definition	9

2.2.2 Input / Output Circuit Diagram	9

2.3 Data Output Type	10

3 Camera Feature	11

3.1 Rolling Mode	11

3.2 HDR Mode	11

3.3 sCMOS Sensor	12

3.4 Offset	12

3.5 Exposure Time Setting	12

3.6 Resolution	12

3.7 ROI	12

3.8 Gain Mode	13

3.10 Gamma	13

3.11 Contrast	13

3.12 Sharpen	13

3.13 Filter	13

4 Camera Operation Modes	13

4.1 Free-Mode	14

4.2 Hardware Trigger Mode	14

4.2.1 Normal Level Trigger Mode	14

4.2.2 Normal Synchronous Readout Trigger Mode	14

4.2.3 Global Exposure Level Trigger Mode	15

4.3 Software Trigger Mode	15

4.4 Trigger Acquisition	15

4.4.1 Close Trigger Acquisition Mode	15

4.4.2 I Open Trigger Acquisition Mode	15

4.4.3 Frames Acquisition	15

5 Application Software	15

5.1 Driver Installation	16

5.2 Application Software Introduction	16

5.3 SamplePro	17

5.3.1 SamplePro Introduction	17

5.3.2 SamplePro Window	17

5.4 Camera Control	17

5.4.1 Resolution	18

5.4.2 Language	18

5.4.3 Exposure Time	18

5.4.4 Frame rate	18

5.4.5 Auto Exposure & Target	18

5.5 Image Configure	18

5.5.1 ROI	19

5.5.2 Auto Level	19

5.5.3 Left & Right	19

5.5.4 Gain Mode	19

5.6 File Operation	19

5.6.1 File Format	19

5.6.2 Parameter Save & Load	19

5.6.3 Capture	20

5.6.4 Code & Record	20

5.7 Other Parameters	20

5.7.1 Flip	20

5.7.2 Output Signal	20

5.7.3 Gamma	20

5.7.4 Contrast	20

5.7.5 Sharpen	21

5.7.6 Filter	21

5.8 Trigger	21

5.8.1 Trigger Mode	21

5.8.1.1 Module Introduction	21

5.8.1.2 API Introduction	21

5.8.1.3  Trigger Function Introduction	22

6 Support	23

- 1 Overview

- 1.1 Manual Instructions

This manual contains basic information, precautions, camera features, and the user guide of software came with the camera.

- 1.2 Naming Rules

The Camera model name includes five parts which are separated by delimiter.

S

401

M/C

G/R

H/U/CL

Shutter

Type

Color

Type

Resolution/

Quantum efficient characteristics and product form

Product Identifier

Output

Type

1               2                3               4                  5

Figure1-1: Camera model instruction

Each part means:

Table 1-1 Camera model instruction

| S | Product Identification |
| --- | --- |
| 400/401 | “400”stands for 4MP camera, ”401”stands for 4MP camera module |
| M/C | M: Monochrome C: Color |
| G/R | G: Global shutter R: Rolling shutter |
| H/U/CL | H:HDMI output U:USB3.0 output CL: Camera Link output |

Note：Product form is determined by mantissa , the mantissa "00" means a camera, if the mantissa "01" means camera model. The table 1-1 takes S-400/401 as an example. Some others use the quantum efficiency characteristic as an identifier，such as S-UV-M-R-U（have response in ultraviolet band）,S-95-M-R-U（quantum efficiency is up to 95 %）.

- 1.3 Packing List

The whole camera set includes the items as below chart:

| Parts | number |
| --- | --- |

Table 1-2 Camera packing list

Please check whether all items are included when get the new cameras. If something is missing or damaged, please contact the after-sale service.

- 1.4 Safety Precautions

Please read the following safety measure to avoid the incorrect use of the camera which could bring damage to the camera or personal property losses.

①Voltage should be within the standard voltage, otherwise it will cause damage to the camera. In serious cases it can cause a fire.

②Do not touch the plug with wet hands when the camera is working.

③Cut off the power supply if the camera is smoking or smell burning, contact us for maintenance.

④Do not disassemble the camera without non-professional workers, avoid damage to the camera by incorrect operations.

⑤Avoid using this camera in the environment such as the temperature is too high or too low, the humidity is very large, very dusty environment, the sunshine is too strong or a place is surrounded by flammable goods and so on.

- 1.5 Dimensions

- 1.5.1 Camera Dimensions

The dimensions of S-UV-M-R-U/HS-95-U3:

Front view                       Rear view

Side view

俯视图

figure 1-2: Dimension of S-UV-M-R-U

The dimensions of HS-200(400)-U3:

Front view                               Rear view

Side view

figure 1-3: Dimension of S-200(400)-M-R-U

- 1.5.2 PCB Dimensions

There are 3 printed circuit boards: Sensor board, FPGA board and USB board. Figure 1-4 shows the S-UV-M-R-U camera's PCB dimensions and Figure 1-5 shows the S-200(400)-M-R-U camera's PCB dimensions.

The S-UV-M-R-U camera's PCB dimensions:

Sensor board                 FPGA board

USB board

figure 1-4: PCB Dimensions of S-UV-M-R-U/HS-95-U3

Sensor board                FPGA board

USB board

figure 1-5: PCB Dimensions of S-200(400)-M-R-U

- 1.6 Camera Features

S-UV-M-R-U camera:

2 inch scientific level CMOS image sensor

4MP, 11μmx11μm pixel size

High ultraviolet quantum efficiency（85%@280 nm）

25 fps @ 4MP

USB 3.0 data transfer

Also provide custom digital interface, support further development.

HS-95-U3

2 inch back-illuminated thinned sCMOS chip

95%@580nm ultra-high quantum efficiency

1.6e-Ultra low read noise

2048×2048 resolution

93dB dynamic range

Support multiple trigger function modes

USB3.0 transmission interface

HS-200(400)-U3：

1 inch or 1.2 inch scientific level CMOS image sensor

2MP or 4MP，6.5μm x 6.5μm pixel size

1.8 e- Readout Noise

0.001 lux minimal illumination

High Quantum Efficiency

64 fps@2MP; 36fps@4MP

USB 3.0 data transfer

Also provide custom digital interface, support further development.

- 1.7  Technical Specifications

Table 1-3 camera technical specifications

| Product | S-UV-M-R-U | HS-95-U3 | HS-400-U3 |
| --- | --- | --- | --- |
| Image Sensor |  |  |  |
| Sensor Type | Scientific CMOS | Scientific sCMOS | Scientific CMOS |
| Sensor Size | 2 inch | 2 inch | 1.2 inch |
| Color Mode | Monochrome | Monochrome | Monochrome |
| Effective Pixels | 4MP | 4MP | 4MP |
| Image Resolution | 2048x2048 | 2048x2048 | 2048x2048 |
| Pixels Size | 11μmx11μm | 11μmx11μm | 6.5μmx6.5μm |
| Camera Parameters |  |  |  |
| Quantum Efficiency | 84.9 %@280nm | 95%@580nm | 72%@625nm |
| Readout Noise | 1.3e- | 1.6e- | 1.8e- |
| Sensitivity Threshold | 3p~ |  | 3p~ |
| Minimal illumination | 0.0001 lux |  | 0.001 lux |
| Full well Capacity | 100ke- | 100 ke | 48ke- |
| Image Acquisition |  |  |  |
| Dynamic Range | 93dB | 93dB | 85dB |
| Shutter Type | Rolling Shutter | Rolling Shutter | Rolling Shutter |
| Exposure Control |  |  |  |
| Exposure Mode | Manual | Manual/auto | Manual/auto |
| Exposure Time | 0.02ms-10 s | 0.02ms-10s | 0.02ms-10s |
| Image Processing |  |  |  |
| Settings | ROI、Gamma、Contrast、External trigger、mirror | ROI、Gamma、Contrast、External trigger、mirror | ROI、Gamma、Contrast、External trigger、mirror |
| Scan Mode | Progressive | Progressive | Progressive |
| A/D | 16bit | 8bit/16bit | 16bit |
| Cooling Mode | N/A | N/A | N/A |
| Preview & Transmission |  |  |  |
| Frame Rate | 24FPS(2048x2048) | 24 FPS (2048x2048) | 36FPS (2048x2048) |
| Data Transmission | USB 3.0 (or parallel digital interface) | USB 3.0(or parallel digital interface) | USB 3.0(or parallel digital interface) |
| Data Cable | USB3.0 cable | 1.8m USB3.0 cable | 1.8m USB3.0 cable |
| Software&compatible system |  |  |  |
| PC software | ISCapture | ISCapture/ ImageJ | ISCapture/ImageJ |
| Compatible system | Windows | Window/Linux/Mac | Window/Linux/Mac |
| Interface |  |  |  |
| Optical Interface | EF mount/ T mount/ C mount | C mount | C mount |
| Data Interface | USB 3.0/5Gb/s | USB3.0/5Gb/s | USB3.0/5Gb/s |
| Power Interface | USB 3.0 power supply or micro USB auxiliary power supply | USB 3.0 power supply or micro USB auxiliary power supply | USB 3.0 power supply or micro USB auxiliary power supply |
| Auxiliary Power Interface | Micro USB 2.0 | Micro USB 2.0 | Micro USB 2.0 |
| Dimension | 63 mm x 63 mm x 63 mm | 63 mm x 63 mm x 63 mm | 50 mm x 50 mm x 62 mm |
| Power Consumer | <2W | <2W | 1.2 W-2.8 W |

- 1.8 Spectral Response

Figure 1-6 shows the quantum efficiency curve of S-UV-M-R-U, the peak QE is 84.9%@280nm.

Figure1-6: S-UV-M-R-U QE curve

Figure 1-7 shows the quantum efficiency curve of HS-95-U3, the peak QE is 95%@580nm.

Figure1-7: HS-95-U3 QE curve

Figure 1-8 shows the quantum efficiency curve of HS-200(400)-U3, the peak QE is 80%@625nm.

Figure1-8: S-200(400)-M-R-U QE curve

- 1.9 Chip Temperature

- 1.9.1 Working Temperature

TBD

- 1.9.2 Relationship between dark current and temperature

The number of electrons generated by the thermal excitation could be suppressed and the dark current could be reduced by lowering the sensor temperature.

Figure1-9: S-UV-M-R- U Dark Current vs. Chip Temperature Relation Curve

- 1.10 Applications

S-UV-M-R-U camera application filed:

UV detection, fingerprint analysis, fire warning, tail flame detection.

S-200(400)-M-R-U camera application filed:

Low light night vision, gun aiming, security monitoring etc.

- 2 Physical Interface and I/O Trigger Connector

Figure2-1 Camera back view

- 2.1 Camera Interface Functions

The table 2-1 below shows interface functions:

Table 2-1 The camera interface functions

| Part name | Function |
| --- | --- |
| TIGGER | External trigger signal input and signal output |
| USB | For the camera power supply and the data transmission |
| LED | Power indicator lights, switch on power will be bright. |
| MICRO USB | Auxiliary power supply(optional) |

- 2.2 The I/O Trigger Pins and Connector Interface Functions

- 2.2.1 The Pins Definition

As shown in figure 1 trigger interface label, the pin definition is in table 2-2 below:

Table 2-2 trigger connector I/O pin definition

| External trigger connector pin number | Pin function and definition |
| --- | --- |
| 1 | TRI_IN（trigger input） |
| 2 | TRI_GND（GND） |
| 3 | TRI_GND（GND） |
| 4 | TRI_OUT0（Output the exposure indicator） |
| 5 | TRI_OUT1（Indicating signal output alone） |
| 6 | NC |

- 2.2.2 Input / Output Circuit Diagram

Figure2-2 Trigger light coupling input circuit diagram

Trigger input to the optical coupling circuit voltage in the range of 2.5~30V.

Figure2-3 Trigger light coupling output circuit diagram

Trigger output to the optical coupling circuit voltage in the range of 3.3~30V.

The relationship between the external trigger input pulse, output indication signal and camera exposure timing is shown in Figure 2-4 (with the rising edge level, the standard trigger mode as an example).

Figure 2-4 Trigger input and output timing diagram

The external trigger level signal through the optocoupler circuit first , and there will be a delay Tiso (6 μs ~ 8 μs delay in rising edge input or 40 μs ~ 60 μs delay in falling edge input, which is determined by the characteristics of the optocoupler circuit)

After the optocoupler delay , the input signal to the camera's internal level has a certain logic delay Tlogic . The  value of Tlogic is 0 ~ 1 minimum exposure unit Tline. The minimum exposure unit of the camera is the exposure time of single line Tline (14μs ~ 21 μs). Namely the logic delay time range is 0 to 21 μs.

Overall, the total delay of the external trigger input to the start of the exposure Tidelay = Tiso + Tlogic.

For three output indication signals, the meanings are as follows:

When the exposure signal is selected, the OUT0 connector will output: the indication level signal of the first line starting exposure to the last line starting exposure.

When the global signal is selected, the OUT1 connector will output: the indication level signal of the last line starting exposure to the first line ending exposure.

When the read signal is selected, the OUT2 connector will be output：the indication level signal of the first line starting readout to the last line starting readout.

These three indicators are always opening by default.

After the three signal output by logic, and then through the optocoupler circuit will have a delay Tiso (6 μs ~ 8 μs delay in rising edge input or 40 μs ~ 60 μs delay in falling edge input)

- 2.3 Data Output Type

A、 The first type: USB3.0 transmission

B、 The second type(customize interface): 40pin, 0.5mm, FPC cable are defined as figure 2-4 shows below.

Consists of parallel TTL digital signal output, I2C control, SPI control, trigger interface.

Figure2-4 Signal pin definition

Note：If the signal pin definition is different from left half of the chart, please according to the table.

- 3 Camera Feature

- 3.1 Rolling Mode

The first pixel line will be exposed when exposure start signal is received in rolling mode. At the end of the exposure time of first row, each pixel value is output immediately and the second row begins to be exposed. Pixel lines take turns to be exposed and readout from the first row to the last row. Pixel data from sensor is output to build image row by row sequentially. Due to the one line time between each row, exposure length is same but exposure timing is different for all lines.

Figure3-1 Row operation in rolling shutter

For its feature of progressive scanning , rolling shutter camera is suitable for capturing static object but not the high speed moving object.

- 3.2 HDR Mode

Image signals are readout from the HG channel and LG channel simultaneously from both bottom and top side of the sensor line by line, a high dynamic range image can be constructed by the low gain image and high gain image for HDR mode.

- 3.3 sCMOS Sensor

Indigo-imaging camera is based on scientific CMOS image sensor, featured with very low readout noise and high frame rate. These features make Indigo-imaging camera an ideal camera for various applications.

- 3.4 Offset

The camera has noise and electron when working in complete dark condition. The noise and electron need to be subtracted because they will be found and considered as valuable signal.

- 3.5 Exposure Time Setting

The value of exposure time is set by software or external trigger.

- 3.6 Resolution

The full resolution of camera is 2048x2048(1x1 bin, bin means pixel binning). The 1x2 bin resolution is 2048x1024, for two of vertical pixels are merged into one pixel. The 2x1 bin resolution is 1024x2048, for two of horizontal pixels are merged into one pixel. The 2x2 bin resolution is 1024x1024, for four of vertical and horizontal pixels are merged into one pixel.

Figure3-2 Pixel binning diagram

- 3.7 ROI

Define an interested area within the scope of the camera sensor resolution to read out the image information and get the image of local area only. The Region of Interest (ROI) control allows users to operate at higher frame rates by placing a lower resolution “window” on the image at any location.

Figure3-3  ROI operation

- 3.8 Gain Mode

The camera has three working modes;

High Gain: High sensitivity mode, which is suit for ultra-low light detection;

Low Gain: High full-well capacity mode, which can detect maximum photon;

HDR: High dynamic range, which can detect ultra-low light and strong light signal simultaneously.

3.9 Signal Output

Output signal is used to output level signal for the user guidance, and the followings show the connection of the trigger interface pin:

Pin ISO 0 output exposure signal

Pin ISO 1 output readout signal

- 3.10 Gamma

Gamma is used to obtain correct reproduction of intensity.

- 3.11 Contrast

Contrast is the difference between the brightness bright and the darkest darks in an image. Higher contrast will make the shadows become darker and the highlights brighter. High contrast will lost more image details.

- 3.12 Sharpen

Sharpness is the contrast on the edge. Sharpening can help to get sharper images and see more image details by increasing the bright and dark lines on edges.

- 3.13 Filter

Each image has noise and the noise has great effect on the quality of image. Filter is used to reduce noise for the noise cannot be undo .

The filter, namely 2D median noise reduction, can remove acnode noise without breaking the image edge features.  It's effective to filter out pulse interference and scan image noise.

- 4 Camera Operation Modes

The camera's operating modes are Free Mode, Hardware-Triggered, and Software-Triggered.

- 4.1 Free-Mode

In free running mode, the camera is reading out and uploading the data continuously. Give a snapping signal, the camera will grab part of uploading data, which may contain incomplete image frame.

- 4.2 Hardware Trigger Mode

There are three types of camera hardware trigger modes:

①Normal level trigger mode

②Normal synchronous readout trigger mode

③Global exposure level trigger mode .

- 4.2.1 Normal Level Trigger Mode

Normal level trigger means that the control of beginning to exposure is set by the falling or raising edge of the signal level and the length of exposure time is determined by the duration of level. This mode does not shoot continuously because of the trigger ready output, commonly used for capturing images of stationary and slow moving objects.

Figure4-1 Normal level trigger

- 4.2.2 Normal Synchronous Readout Trigger Mode

Synchronous trigger read out is the exposure of each frame of the camera are subject to external trigger source control, that is synchronous read trigger time is determined by the external trigger, the camera's read and the next exposure is synchronized.

Figure4-2 Normal synchronous readout trigger

- 4.2.3 Global Exposure Level Trigger Mode

Global exposure level trigger has four characteristics:

1) The camera is pre-exposed, but not scanned.

2) The light source triggers the camera to begin exposure, and the lighting phase must ensure that all lines are in the exposure phase.

3) Exposure time is determined by the level width, and the output data consists of 2 successive frames.

4) It is not a continuous shooting mode.

Figure4-3 Global exposure level trigger

- 4.3 Software Trigger Mode

Set the camera in the software mode, give the camera a snap signal, the camera will capture an image with a full frame from the memory block.

- 4.4 Trigger Acquisition

- 4.4.1 Close Trigger Acquisition Mode

When the trigger acquisition mode is set to off, there is no need to request to trigger the acquisition signal, the camera will generate by its own internal.

- 4.4.2 I Open Trigger Acquisition Mode

When Trigger Acquisition Mode is set to open, the camera is waiting for the acquisition to begin. A start trigger signal must be requested to allow the camera to exit the waiting and begin acquisition state and enter the acquisition state waiting frame start trigger before the camera can begin frame acquisition.

- 4.4.3 Frames Acquisition

If the trigger acquisition mode is set to on, you need to set the number of frames to capture. The number of acquired frames can be set between 1 and 255.

- 5 Application Software

SamplePro is excellent software for camera control, image acquisition and archiving of images in various file formats. This chapter provides a detailed description of all SamplePro functions.

- 5.1 Driver Installation

To use the Indigo-imaging camera, you need to install the camera driver and SimplePro application software first.

Step1. Find the file “Indigo-imaging Camera Driver " Setup.exe" from the CD, double-click on it to start the installation and hit “Next”.

Step2. Follow the steps to finish the installation.

NOTE：Please go to the Device Manager to check if the driver was installed properly. When the driver is installed correctly, there is no yellow mark with the camera under the Imaging Device in Device Manager.

- 5.2 Application Software Introduction

The SamplePro windows application software enables you to control every camera parameter or setting. Image can be displayed on a monitor and may be stored. The table below lists the key functions of SamplePro.

| Application Software |  | SamplePro |
| --- | --- | --- |
| Function Menu |  | Introduction |
| Camera Control | Preview | Start the program and display the live images in the window |
|  | Resolution | Select resolution for live image and capturing |
|  | Language | Chinese or English |
|  | Exposure Time | Adjust the exposure time manually |
|  | Frame rate | Display the current frame rate |
|  | Auto Exposure & Target | Check the "Auto Exposure" check box to adjust exposure time automatically |
| Image Configure | ROI | Place a lower resolution “window” on the image at any location |
|  | Auto Level | Turn on/off the auto level function |
|  | Left | Lower limit of the level displayed |
|  | Right | Upper limit of the level displayed |
|  | Gain Mode | Select the working gain mode：high gain, low gain, HDR |
| File Operation | File Save | Select the format of images：jpg, tiff, png or raw |
|  | Parameter Save & Load | Save the current parameter settings or load the saved parameter settings |
|  | Capture | Capture the image |
|  | Code & Record | start/ stop video recording |
| Other Parameters | Flip | Check and make the image flip |
|  | Output Signal | Turn on/off signal output |
|  | Gama | Set the gamma value |
|  | Contrast | Set the contrast value |
|  | Sharpen | Set the sharpness value |
|  | Filter | Turn on/off the filter function |
| Trigger | Hardware | Select the external hardware trigger mode |
|  | Exposure | Select exposure time mode |
|  | Edge | Select activation edge |
|  | Frames | Adjust the frames per capturing |
|  | Delays | Adjust the delay time to trigger exposure |
|  | Apply | Make the trigger mode that is choose working |
|  | Software-On | Turn on/off the software trigger function |

Table 5-1 SamplePro function list

- 5.3 SamplePro

- 5.3.1 SamplePro Introduction

SamplePro is used to obtain and process pixel date from Indigo-imaging camera. It consists of live image window and picture processing module. User can control the camera easily by SamplePro.

- 5.3.2 SamplePro Window

Figure5-1 Software window

Figure5-1 shows the interface of SamplePro, the left is menu bar and the right is live image window.

- 5.4 Camera Control

Figure5-2 Camera control menu

- 5.4.1 Resolution

Adjust the resolution of image.

- 5.4.2 Language

Select the displayed language for software: Chinese (simplified), English.

- 5.4.3 Exposure Time

Adjust the exposure time manually. Key the exposure time in the edit box directly, then click "OK" to confirm it.

- 5.4.4 Frame rate

Display the current frame rate.

- 5.4.5 Auto Exposure & Target

Check [Auto Exposure] checkbox, software start to adjust the exposure time automatically to get proper brightness of live image.

Auto exposure target value: Set the reference exposure time for auto exposure adjustment. The default 110 is recommended. It helps the auto exposure to find the proper exposure time faster. For example, if the imaging target is quite bright, a lower target value is set, it will indicate a shorter exposure time to the software.

- 5.5 Image Configure

Figure5-3 Image configure

- 5.5.1 ROI

It is used to set the interested areas of image, with the upper left corner as the origin of coordinates.

- 5.5.2 Auto Level

Check the "Auto Level" checkbox: adjust the live image levels automatically.

- 5.5.3 Left & Right

Left: It is used to set the black point or shadow values. Move the Left Point Slider towards right to darken the shadow areas in the image. It will reveal bright area information.

Right: It is used to set the white point or shadow values. Move the Right Point Slider towards left to brighten the highlights in the image. It is able to reveal some information in dark area.

- 5.5.4 Gain Mode

The camera's working modes are divided into the following 3 categories: high gain, low gain and HDR.

- 5.6 File Operation

Figure5-4 File operation menu

- 5.6.1 File Format

Four file formats are available: TIF,PNG,JPG and RAW. JPG is the default format.

NOTE: Raw and tiff image file contains minimally processed data from the camera. It needs to be read in some special software for example Photoshop, ImageJ etc.

- 5.6.2 Parameter Save & Load

Save: save the current parameters information to item A, B or C .

Load: load the setting parameters to current image from item A, B or C.

Item A/B/C is used to keep the parameters setting, just like the text file.

- 5.6.3 Capture

Capture the image with current parameter settings.

- 5.6.4 Code & Record

Click "Record", start/ stop video recording.

Click [Code] to list all the available video compressors on the PC.

NOTE: The video taken without any compression will be very large size.

- 5.7 Other Parameters

Figure5-5 Other parameters menu

- 5.7.1 Flip

Check the "HFlip" checkbox to make the image flip horizontally.

Check the "VFlip" checkbox to make the image flip vertically.

- 5.7.2 Output Signal

Check the "Exposure", one trigger interface will output : the indication level signal of the first line starting exposure to the last line starting exposure.

Check the "Readout", another trigger interface will output: the indication level signal of the first line starting readout to the last line starting readout.

- 5.7.3 Gamma

Adjust image gamma, the range is 0.2 to 1.5 and 1.0 is default. Gamma is used to obtain correct reproduction of intensity.

- 5.7.4 Contrast

Adjust Contrast. The range is 0-255 and 128 is default. Contrast is the difference between the brightness brights and the darkest darks in a image. Increase the contrast, the shadows become darker and the highlights brighter. Decrease the contrast, the highlights grow dim and the dark areas lighten up.

- 5.7.5 Sharpen

Adjust the image sharpness.

- 5.7.6 Filter

Check the "Filter" checkbox to open the filter function. Opening is default setting.

- 5.8 Trigger

- 5.8.1 Trigger Mode

Figure5-6 Trigger menu

- 5.8.1.1 Module Introduction

Camera capture modes are divided into the following 2 categories:

Sequence mode (stream mode): it is used to capture continuous image data.

Trigger mode: the camera captures the images by the external signals.

Trigger Mode:

1. Standard Mode, the parameters of exp/edge/delay/frames are supported.

2. Synchronization Mode, namely the synchronous picture grabbing, only the parameter of edge is supported.

3. Global Mode, used to control the CMOS camera of shutter exposure mode to achieve global exposure mode like CCD, generally used for the scenes of controllable light source.

4. Software Trigger Mode: send commands to simulate trigger signal by software.

- 5.8.1.2 API Introduction

1Typedefs enum { DEVICE_STREAM = 0x0001, DEVICE_HW_STANDARD = 0x0002, DEVICE_HW_SYNC, DEVICE_HW_GLOBAL, DEVICE_SW_TRI, // software trigger DEVICE_RESEVER , }; 2 function declaration SetTrigger(int mode, int exp, int edge, int delay, int frames, const char *devName); GetTrigger(int *pMode, int *pExp, int *pEdge, int *pDelay, int *pFrame, const char *devName); SetGpio(int index, bool state, const char *devName); 3 parameters explanation mode: 0~3,hardware[send command by hardware]，4 software [send command by SetGpio()] datamode,0 standard steam mode [hardware],all of the trigger parameters are not supported. datamode,1 standard trigger mode [hardware],all of the trigger parameters are supported. datamode,2 Synchronization trigger mode [hardware],only edgetype is supported. datamode,3 Global trigger mode [hardware],exptype and edgetype are supported. datamode,4 Synchronization trigger mode [software], all of the trigger parameters are not supported, send command by SetGpio(). exptype: exposure type 0：timed，1：level width edgetype: 0:rising edge,1:falling edge delaytime：the delay time to exposure, the unit is ms. frames： The number of captured frames

- 5.8.1.3  Trigger Function Introduction

0) Closing ：

Call SetTrigger(0, x, x, x, x,pName) into DEVICE_STREAM mode , means trigger mode is closed. The camera captures continuous image data.

When choosing "Standard"," Synchronization" or "Global", the camera switches to trigger mode. The camera only captures image data after receiving a trigger signal.

Attention：the camera will capture images if input level signal is received even then the trigger mode is closed. This is a special design in the camera.

- Standard Mode：

Call SetTrigger(1, x, x, x, x,pName) into DEVICE_HW_STANDARD mode;

Standard Mode consists of the following configurations:

A. Exposure time:

Timed: after receiving the trigger signal, it is determined by the exposure time set by the software.

Trigger Width: after receiving the trigger signal, the exposure time is determined by the level width.

B. Activation Edge

Rising Edge: means the trigger signal is rising edge valid

Falling Edge: means the trigger signal is falling edge valid

C. Number of frames

It means after receiving a trigger signal, the exposure time of each image is the same no matter how many images are shot, which depends on the software settings. (When choosing level width, the function is invalid.)

D. Delay

It means after receiving a trigger signal, the delay time of the desired delay time can be set to trigger the camera exposure .

- Synchronization Mode：

Call SetTrigger(2, x, x, x, x,pName) into DEVICE_HW_SYNC mode;

Namely the synchronous picture grabbing, first trigger to start, and the second trigger to output synchronization image.

Synchronization mode has one configuration only, for the exposure time is set by the input level width. Activation Edge：

Rising Edge: means the trigger signal is rising edge valid.

Falling Edge: means the trigger signal is falling edge valid.

3）Global Mode：

Call SetTrigger(3, x, x, x, x,pName) into DEVICE_HW_GLOBAL mode;

A method that uses the rolling shutter readout mode with a slightly expanded exposure time, resulting in a slice of time during which all the lines of the sensor are being exposed simultaneously

Global mode consists of 2 configurations:

A. Exposure time:

Timed: means the exposure time is set by the software

Trigger Width: means the exposure time is set by the input level width

B. Activation Edge

Rising Edge: means the trigger signal is rising edge valid

Falling Edge: means the trigger signal is falling edge valid

4）Software Trigger：

Simulate trigger signal via the software command.

Call SetTrigger(4, x, x, x, x,pName) into DEVICE_SW_TRI mode.

Click "snap" to capture images by sending command to camera . (send SetGpio(0, true,camName));

NOTE：Click "Apply" bottom to confirm your setting when the parameters of mode/exp/edge/delay/frames is changed.

- 6 Support

You can also get support in the following ways:

- Web Support - Visit www.indigo-imaging.com for documentation and technical support.

- Hotline Support - Contact us via hotline 086-0591-28055076

- E-mail support - Send feedback to support@indigo-imaging.com, our technical staff will reply in time.